## Slide 1: scikit-rf: an open-source Python package for RF network simulation, analysis and calibration
The scikit-rf team https://github.com/scikit-rf/scikit-rf/graphs/contributors
http://scikit-rf.org/

## Slide 2: RF Engineering Typical Problems
Reading VNA Touchstone Files (.sNp)
  2-ports, 4-ports or 8-ports: different formatting!
Extract and Plot Network Parameters
  What is the formula to convert S to Z?
Compare, Combine Networks
  Incertitude of a Series of Measurements?
  Cascading or Deembedding
Calibration
  Problem with Calibration = Unusable Measurements!
  Tiered Calibration (Ex: 2-port measured with a 1-port setup)
I don’t have the license for the XXX software/toolbox!
©INGRAM PUBLISHING
Etc…

## Slide 3: scikit-rf: an open-source package for RF engineering
A modern, object-oriented library for network analysis and calibration
  Designed to make RF/microwave engineering both robust and approachable
  Created in 2009 by Alex Arsenovic and continually developed since
    By more than 50 volunteers on GitHub (https://github.com/scikit-rf/scikit-rf)
    Distributed with BSD open-source license: benefits to the needs of private industry and gov. contracts
    >23 000 lines of code, estimated 6 years of effort
  Used by several companies, universities and research groups in the world

## Slide 4: Introduction to scikit-rf
The Network Class

## Slide 5: Network : the scikit-rf base class
Any RF Network is fully described by its S-parameters
  Represented in scikit-rf with a Network object
    Created Manually from S-parameters (or Z, Y, etc.)
    Or imported from a Touchstone file (.sNp)
    >>> import skrf
    >>> ring_slot = skrf.Network('data/ring slot.s2p')
  The Short Description of a Network object gives
    The number of port(s),
    The frequency range and number of points of the network parameters,
    The Characteristic impedances of its ports (which can be frequency-dependant) :
    >>> print(ring_slot)
    2-Port Network: 'ring slot', 75.0-110.0 GHz, 201 pts, z0=[50.+0.j 50.+0.j]

## Slide 6: Network : the scikit-rf base class
Network parameters and associated projections are easily accessible
  For example, for the most used ones:
    >>> ring_slot.frequency # frequency band properties
    >>> ring_slot.z0 # ports characteristic impedances
    >>> ring_slot.s # S-parameters. Dimension is (nb_f x nb_port x nb_port)
    >>> ring_slot.z # Z-parameters. Dimension is (nb_f x nb_port x nb_port)
    Also available: .y, .a (ABCD), .t (transfer), .h (hybrid)
  Several Available Projections: complex, real, imag, magnitude, dB, phase (rad/deg) :
    >>> ring_slot.s_db # S-parameters magnitude in dB
    >>> ring_slot.z_deg # Z-parameters phase in degree
  Extract Data for Sub-Frequency Bands, using frequency indices… or using « natural » language:
    >>> ring_slot[100:200]
    2-Port Network: 'ring slot', 92.5-109.825 GHz, 100 pts, z0=[50.+0.j 50.+0.j]
    >>> ring_slot['80-100GHz']
    2-Port Network: 'ring slot', 80.075-100.025 GHz, 115 pts, z0=[50.+0.j 50.+0.j]

## Slide 7: Graphical Representations
    >>> ring_slot.plot_s_db()
    >>> ring_slot.plot_s_smith()

## Slide 8: Building RF Circuits

## Slide 9: Cascading Networks
scikit-rf allows to connect N-ports Networks easily
  Cascading between two Networks is performed using the ** Python Operator
  Cascade of a 2-port with a 1-port
  Cascade of two 2-port
  Cascade of two 2N-port

## Slide 10: Cascading and Simple Deembedding
  Cascade of a transmission line with a short-circuit:
    >>> short = skrf.data.wr2p2_short # WR2.2 (1-port, 330-500 GHz) short
    >>> line = skrf.data.wr2p2_line # WR2.2 (2-port, 330-500 GHz) line
    >>> delayshort = line ** short
  Deembedding of a Network of known S-parameters
    >>> short_2 = line.inv ** delayshort # deembedding

## Slide 11: scikit-rf can be used to build advanced Circuits
N-ports Circuit Creation
  From the assembly of several Networks
  Calculation of S-parameters (at ports or internally), V/I at nodes

## Slide 12: Another Example of Advanced Circuit
Combination of Lumped Components and Full-Wave Results
Hillairet, J. RF network analysis of the WEST ICRH antenna with the open-source python scikit-RF package. AIP Conference Proceedings 2254, 070010 (2020).
Touchstone
from HFSS
Touchstone
from HFSS
Touchstone
from HFSS

## Slide 13: Time Domain Analysis

## Slide 14: Time Domain and Gating
Suppression of parasitic reflections
  Waveguide alignment can be problematic at very high frequency (>100 GHz)
    Generate parasitic reflection
    S-parameters are measured in the frequency domain, but can be analyzed in time domain
    Parasitic phenomenon can be suppressed by filtering the time response (time-gating)
WR-2.2/UG-387 325 – 500 GHz Gain Horn Antennas and straight waveguide (© MI-WAVE)
0.022” x 0.011”
(0.56 x 0.28 mm)
Notes: CNC milling offers ±2.5−μm tolerances at best

## Slide 15: Time Domain and Gating
Suppression of parasitic reflections
  Example: isolation of a particular response in a signal affected by multi-reflections
    >>> probe = skrf.Network('probe.s2p') # CPW probe @ 500-750 GHz
    >>> s11 = probe.s11
    >>> s11_gated = s11.time_gate(center=0, span=.2)# gate 1st largest reflec

## Slide 16: Measurement Sets and Statistics

## Slide 17: NetworkSet: A Set of Measurements/Networks
scikit-rf eases series of measurements analysis with the NetworkSet class
  Reading multiple .sNp files at once
  Deduce statistical properties (mean, standard-deviation, covariance, etc.)
    >>> from skrf.data import ro_1, ro_2, ro_3 # ex of 3 consecutives measurements
    >>> ntwk_list = skrf.NetworkSet([ro_1, ro_2, ro_3])
    >>> ntwk_list.plot_uncertainty_bounds_s_db() # 3-sigma uncertainty

## Slide 18: NetworkSet: Named Parameters and Interpolation
NetworkSet
scikit-rf can read data set files such as CITI (.cti) or GMDIF (.mdif)
  Content of these files is converted into a NetworkSet
    Includes all the defined named parameters (ex: V=0, 5, 10; A=1)
  Get data subset from named parameters
    Ex: all measurements made for V = [0 to 10] and A = 1
Interpolation of data set
  Interpolation of network parameters for a data set
    Ex: Interpolation of S-parameters for intermediate values of V
V=0, A=1
Network
V=5, A=1
Network
V=10, A=1
Network
V=2, A=2
Network
V=2.6, A=2
Network
V=0, A=1
Network
V=5, A=1
Network
V=10, A=1
Network
NetworkSet
V=4, A=2
Network
sel
interp

## Slide 19: Calibration, Deembedding

## Slide 20: Off-Line Calibration in scikit-rf
« Off-line » calibration: flexibility and raw data preservation
  Several calibration algorithms are available
    1-port (SOL, SDDL, PHN, etc.)
    2-port (SOLT, (multiline) TRL, Unknown-Thru, LRM, LRRM, 8, 12 ou 16-terms generic models, etc.)
  Example: 2-port TRL calibration:
    >>> measured = [T,R,L]
    >>> trl = TRL(measured)
    >>> dut_corrected = trl.apply_cal(dut_raw)
Notes: o measure the switch terms manually, you can define custom traces as ratio’s of receivers and save the data to touchstone files. Assume ports are labeled 1 and 2, and incident waves are a’s and reflected waves are b’s. Then, the
forward switch term == a2/b2 with source port 1
reverse switch term == a1/b1 with source port 2

## Slide 21: Deembedding
Deembedding: removing the effects a test fixture can have on a DUT
raw_dut

## Slide 22: Deembedding
Deembedding: removing the effects a test fixture can have on a DUT
  OpenShort Deembedding: remove the DUT and measure the fixture with 2 dummy setups:
    >>> dm = OpenShort(dummy_open=open_nw, dummy_short=short_nw)
    >>> dut = dm.deembed(raw_dut)
short_nw
open_nw

## Slide 23: Deembedding
Deembedding: removing the effects a test fixture can have on a DUT
  At frequencies >10 GHz, fringe capacitance and the parasitic inductance cause errors
  Several methods using a through (thru) dummy are implemented in scikit-rf:
    Split-Pi, Split-Tee, Admittance-Cancel (Mangan’s method), Impedance-Cancel

## Slide 24: Summary

## Slide 25: scikit-rf: an open-source package for RF engineering
A modern, object-oriented Python package for network analysis and calibration
  Designed to make RF/microwave engineering both robust and approachable (and free!)
  Eases reading/writing RF data files (Touchstone, CITI, GMDIF, …)
  Eases graphical representations
    dB, amplitude, phase, Smith chart, incertitude, etc.
  S-parameters Transformations
    Conversions (Z, Y, ABCD, etc.),
    Frequency subset, concatenations, sub-network,
    Interpolations between frequency points or sets of data
    Time Domain, Gating
  Off-Line Calibration and deembedding
    1-port or 2-port (SOL(T), SDDL, TRL, LR(R)M, etc.)
    Advanced Deembedding methods
  But Also!
    SPICE circuit modelling (vector fitting)
    Direct communication with VNAs
    Mixed-mode conversions
    Transmission Line modelling
    Etc…
  Acknowledgments
  The project would not exist without the several benevolent contributors!

## Slide 26: Questions or comments?
Website and Documentation
  http://scikit-rf.org/
Ask the Community
  Email: scikit-rf mailing List: http://groups.google.com/group/scikit-rf
  Chat: scikit-rf have channels on both Slack and Matrix
Contributing to the scikit-rf project
  There is multiple ways to contribute to an open-source project like scikit-rf!
    Sponsoring the project
    Contributing to the code
    Contributing to the documentation

## Slide 27
A.Arsenovic, J.Hillairet, J.Anderson et al., "scikit-rf: An Open Source Python Package for Microwave Network Creation, Analysis, and Calibration [Speaker’s Corner]," in IEEE Microwave Magazine, vol. 23, no. 1, pp. 98-105, Jan. 2022, doi: 10.1109/MMM.2021.3117139.
https://ieeexplore.ieee.org/document/9632487
Acknowledgments
scikit-rf would not be possible without the feedback, bug fixes, and new features contributed by its user community.
The complete list of contributors can be seen here:
https://github.com/scikit-rf/scikit-rf/graphs/contributors

## Slide 28: Spare Slides

## Slide 29: scikit-rf: Installation
If you have installed Python, you can install scikit-rf from:
  pip (PyPi):
    python -m pip install scikit-rf
  conda (Anaconda):
    conda install -c conda-forge scikit-rf

## Slide 30: Main application domains of scikit-rf (2021 survey)